# PeakSimple WITS Output — Implementation Guide

This guide explains how to configure PeakSimple to output gas chromatograph

results as WITS Level 0 records over a serial (COM) port, for consumption by

a rig data system (e.g. a Pason EDR unit).

WITS output was introduced in PeakSimple version 4.34. Make sure you are

running 4.34 or later before proceeding.

―――――――――――――――――――――――――――――――――――――――――――――――――――――――

# 1. Overview

Once enabled, PeakSimple will:

Send a WITS keep-alive packet every 20 seconds, so the receiving

equipment knows the link is live.

At the end of each run (during Postrun), send one WITS data packet for

every component that has been assigned a WITS code.

There is currently no dedicated user-interface screen for this feature — the

serial port is enabled and set by hand-editing the control file, and the WITS

code/value mapping is set per-component from the existing Component Details

screen. This guide covers both steps.

―――――――――――――――――――――――――――――――――――――――――――――――――――――――

# 2. Step 1 — Enable the WITS output port

The WITS port number is stored in the control file (the `.CON` file loaded

by PeakSimple, e.g. `DEFAULT.CON`) as the entry:

<WITS PORT>=n

`n` is the number of the COM port to use (e.g. `5` for COM5).

`n = 0` (the default) disables WITS output entirely.

To set it, since there is no on-screen control for this field:

1. Close PeakSimple.

2. Open the control file in Notepad (or a similar plain-text editor).

3. Find (or add) the `<WITS PORT>=` line and set it to the desired COM port

number.

4. Save the file, then reopen PeakSimple so the change takes effect.

**Note:** If you save any control file from PeakSimple 4.34 without

having set this manually first, it will write out `<WITS PORT>=0`

(disabled), which you can then edit in Notepad as above. Changing the

value while PeakSimple is open may not "stick" reliably — closing and

reopening PeakSimple after editing is the most dependable way to apply it.

Serial format: the COM port is fixed at 9600 baud, 8 data bits, no

parity, 1 stop bit (9600,8,N,1). Configure the receiving instrument to

match.

―――――――――――――――――――――――――――――――――――――――――――――――――――――――

# 3. Step 2 — Assign WITS codes to components

WITS codes are attached to components using the existing User

Calculation fields (Calc1–Calc8), so no new dialog is involved:

1. Go to Channel *n* Components → Component Details for the peak/component

you want to export.

2. Open User Calculations.

3. In Calc1, enter the WITS code for that component, e.g.:

Calc1: 1215

This must be a numeric code in the range 1200–1299 — this is the

range PeakSimple recognizes as valid GC data for WITS export. Anything

outside this range will not be sent.

4. In Calc2, enter the field whose value should be sent alongside the

code. This is a normal user-calculation formula, so it can reference any

result field, most commonly:

Calc2: [Area]

or

Calc2: [External]

5. Repeat for each component you want exported, using a different WITS code

for each.

Each qualifying component now has a code (Calc1) and a value expression

(Calc2). The exported value is written with 2 decimal places.

―――――――――――――――――――――――――――――――――――――――――――――――――――――――

# 4. What gets sent, and when

Keep-alive packets are sent automatically every 20 seconds whenever

WITS output is enabled, independent of any run.

Data packets are sent only as part of Postrun processing, i.e.

after a run completes — one packet per component with a valid WITS code

(1200–1299).

Each packet is delimited with `&&` at the start and `!!` at the end.

Keep-alive packets contain a placeholder payload such as `0111-9999`.

Data packets contain the WITS code immediately followed by the value, with

no space between them and 2 decimal places on the value, for example:

1215939.63
  1216213.77

(WITS code `1215` with value `939.63`, code `1216` with value `213.77`.)

―――――――――――――――――――――――――――――――――――――――――――――――――――――――

# 5. Testing the output

You do not need to wait for a full run to test this:

Press Ctrl-Alt-P to immediately trigger all Postrun actions

(including the WITS export) without waiting for a run to finish.

To see the raw output, use a free serial port monitoring utility (for

example, "Free Serial Analyzer" / "Free Device Monitoring Studio") on a

PC connected to the configured COM port. Connect the monitor to the WITS

COM port and you should see the `&&…!!` packets described above appear,

including the periodic keep-alives and, after Ctrl-Alt-P or a real run,

the data packets for your configured components.

―――――――――――――――――――――――――――――――――――――――――――――――――――――――

# 6. Known limitations (as of v4.34)

There is no in-app UI yet for setting the WITS COM port — it must be

edited directly in the control file as described in Step 1.

There is no in-app UI for reviewing which components/codes are currently

configured for WITS export beyond opening each component's User

Calculations screen individually.

WITS codes must fall within 1200–1299; components with codes outside this

range are silently excluded from export.

―――――――――――――――――――――――――――――――――――――――――――――――――――――――

# 7. Quick checklist for a new installation

1. Confirm PeakSimple is version 4.34 or later.

2. Edit the control file in Notepad and set `<WITS PORT>=n` for your COM

port; save and reopen PeakSimple.

3. Configure the receiving instrument's serial port for 9600, 8, N, 1.

4. For each component to export, set Calc1 to a WITS code (1200–1299) and

Calc2 to the result field to send (e.g. `[Area]` or `[External]`).

5. Press Ctrl-Alt-P to force a test export, and confirm the packets with a

serial monitor tool.

6. Once confirmed, run normally — the configured components will be

exported automatically at the end of every run, alongside the periodic

keep-alive signal.